EXPERIMENT 10

#### CHARACTERISTICS OF THE PHOTDIODE

#### OBJECTIVES:

The objective of this experiment is to understand and measure some of the many important characteristics of the photodiode. The photodiode characteristics that will be studied in this experiment are responsivity, I-V characteristics and linearity.

#### EQUIPMENT REQUIRED

1- Optical power meter: INFOS, Model # M100.

2- HeNe laser: Coherent, Model # 31-2090-000.

3- Optical bench (¼ m).

4- Horizontal and vertical stages (Used to carry and align the HeNe laser).

5- Laboratory jack.

6- Multimeter.

7- Power supply.

8- Decade resistor box.

9-  resistor.

10- Medium area silicon photodiode (RS Part Number: 651-995).

11- Polarizer/Analyzer pair.

12- Large carton box with a pinhole.

PRELAB ASSIGNMENT

Read the introduction to this experiment, before you attend the laboratory.

INTRODUCTION:

The photodiode is an essential element in fiber optics communication systems. It is similar to the ordinary semiconductor diode, with one exception: it responds to light. The symbol used for the photodiode is the same as the symbol of the ordinary diode, as seen in Figure 1. The diode voltage  is defined with + polarity on the p terminal and the diode current  is defined to flow through the diode from the p to the n terminal, which is identical to the standard definitions of the voltage and current of the ordinary diode.

Figure 1: Photodiode Symbol, Showing Definition of the Photodiode Current and Voltage.

The simplest way to use a photodiode for the purpose of light detection is to leave its terminals open-circuited, as shown in Figure 2(a). This is an example of the so called photovolatic mode of diode operation. In this simple circuit, the photodiode conducts no current (because of the open circuit). When the incident light power  is zero,  is also zero. However, in this circuit, when , a small (at most a fraction of a volt), but measurable voltage  appears across the photodiode. This voltage is termed photovoltage, because it is caused by the incident light power and it disappears when the incident light power becomes zero. The value of  depends on  as well as on the incident light wavelength . An important example of the application of the photovolatic mode is in solar cells.

(a)                                                             (b)

Figure 2: (a) Open Circuited and (b) Short Circuited Photodiode Circuits.

The circuit shown in Figure 2 (b) is an example of the photoconductive mode of diode operation. In this circuit, because the photodiode is short circuited, the diode voltage  is forced to be zero. The diode produces electric current  in response to the incident light. The current  is called the photocurrent, because it caused by the incident light power and it disappears when the incident light power is zero. This photocurrent is proportional to the incident light power . In general, the higher the light power is, the higher the photocurrent  is. We can then view the photodiode in this circuit as a device that coverts light energy to electric current. The actual direction of the photocurrent is opposite to the one shown in Figure 2 (b). Thus, when the photodiode operates in the photoconductive mode,  is always a negative quantity. Let us then define the current . Because  is positive, it is easier to use in our forthcoming discussion and the experimental work to be done.

For fiber optics communication, the photoconductive mode of operation is used to covert communication signals carried by light to signals carried by electric current. Thus, in this laboratory experiment, we will concentrate on the photoconductive mode of photodiode operation. The relationship between the incident power  and the induced current  is given by:

(1)

Where  is the photodiode responsivity (see Appendix C). The variation of  with  for the medium area photodiode used in this experiment is shown in Figure 3.

Figure 3: Responsivity Versus Wavelength for the Medium Area Photodiode.

For normal operation, the photodiode is reversed biased (i.e. ). The responsivity  of the diode is largely independent of the bias voltage, as long as the diode is reversed biased. When the diode has zero bias (i.e. ),  decreases slightly below the reversed bias value. This effect can be seen clearly seen in Figure 3, where the 12 V indicated in the figure refers to the value of the reverse bias voltage. However, when the diode is forward biased,  begin to drop, reaching negligible values when the forward is sufficiently high. It is thus important to keep in mind that the diode responsivity decreases drastically when the diode is forward biased. This important feature of the photodiode will be observed in this experiment.

Now, consider the circuit shown in Figure 4, which a simple and practical circuit for light detection. The supply voltage  is kept positive. Using Kirchhoff’s voltage law and Ohm’s law, we have:

(2)

(3)

Let us assume that the incident optical power is fixed at a certain level . When the load resistance  equals zero, . Since , then  and thus in this case the diode is reversed biased.  Since  is fixed, we also expect  to also be fixed, as long as the diode remains reverse biased. When  increases, the voltage  also increases. Thus we can continue to increase  by simply increasing  as long as the photodiode remains reversed biased. However, at some point,  will exceed  causing the diode to be forward biased and becoming less and less responsive to the incident power (i.e.  starts to decrease, causing  to decrease). In short, the maximum obtainable value of the load voltage  is slightly larger than the value of . Thus, for this circuit, the following relationship applies:

(4)

This important relationship will be verified in this experiment.

Figure 4: Practical Simple Circuit for Light Detection.

#### PROCEDURE:

Part A: Using the Photodiode in Light Power Measurement and Diode Responsivity.

1- Connect the circuit shown in Figure 5.

2- Turn the HeNe laser on and keep it approximately 1m from the photodiode.

3- Align the laser so that the center of the laser beam is approximately normal and in the middle of the light sensitive of the photodiode. Notice that laser beam is completely within the light sensitive area of the diode, so that all the laser power is sensed by the diode.

4- Use an object to block the laser beam and read the ammeter reading. Record the value in table 1.

5- Remove the object and read the ammeter reading. Verify that the induced current actually flows through the diode from the n terminal to the p terminal, not the other way around. Record the results in table 1.

6- Use Figure 3, equation (1) and the ammeter reading of step (5) to estimate the output power of the HeNe laser. Record the value in table 2.

7- Compare your result with the HeNe Laser specifications. Discuss the results and draw some conclusions.

Figure 5: Setup Used for Measuring the Power Emitted by a HeNe Laser Source.

| Condition | (mA) | Actual Direction of |
| --- | --- | --- |
| Laser Blocked |  | From n to p |
| Laser Unblocked |  | From p to n |

Table 1: Photocurrent Current Measurement for Part A of the Experiment.

| in (A/W) | (mA) | HeNe Laser Power (mW) (Experimental) | HeNe Laser Power (mW) (Manufacturer’s Specification) |
| --- | --- | --- | --- |

Table 2: Photodiode Responsivity and Estimated Power of the HeNe Laser.

Part B: Practical Circuit for Light Detection and Photodiode IV Characteristics.

1- Connect the circuit shown in Figure 6 and repeat steps 1) and 2) of part A.

2- Set  to 0 Volts [Disconnect the leads from the power and connect them together to produce a short circuit, which is equivalent to . Turning the power supply off or setting the power supply voltage to zero does not work well in this case].

3- Set the decade resistance  to  and then measure  and record its value in table 3.

4- Repeat step 3 for , , until the third column of table 3 is filled. To simulate , remove the lead that connects the decade box to the  resistor.

5- Calculate the photodiode current and voltage and record them in table 3.

6- Connect the circuit to the power supply and set  to 5 Volts [You should use a digital volt meter to measure  to set it precisely to 5V. [Do not rely on the power supply meter for this measurement]. Record the precise value of  in the top row of table 4.

7- Set the decade resistance  to , then measure  and record its value in table 4.

8- Repeat step 8 for , , until the third column of table 4 is filled.

9- Calculate the photodiode current and voltage and record them in table 4.

10- Repeats steps 6 through 9, for  (precise) and record the results in table 5.

Figure 6: Circuit Diagram for Part B.

| 0 | 160 |  |  |  |
| --- | --- | --- | --- | --- |
| 100 | 260 |  |  |  |
| 200 | 360 |  |  |  |
| 300 | 460 |  |  |  |
| 400 | 560 |  |  |  |
| 500 | 660 |  |  |  |
| 600 | 760 |  |  |  |
| 700 | 860 |  |  |  |
| 800 | 960 |  |  |  |
| 900 | 1060 |  |  |  |
| 1000 | 1160 |  |  |  |
| 10000 | 10160 |  |  |  |

Table 3: Measured Load Voltage Versus Load Resistance for .

| (Measured) = Volts |
| --- |
| 0 | 160 |  |  |  |
| 100 | 260 |  |  |  |
| 200 | 360 |  |  |  |
| 300 | 460 |  |  |  |
| 400 | 560 |  |  |  |
| 500 | 660 |  |  |  |
| 600 | 760 |  |  |  |
| 700 | 860 |  |  |  |
| 800 | 960 |  |  |  |
| 900 | 1060 |  |  |  |
| 1000 | 1160 |  |  |  |
| 1100 | 1260 |  |  |  |
| 1200 | 1360 |  |  |  |
| 1300 | 1460 |  |  |  |
| 1400 | 1560 |  |  |  |
| 1500 | 1660 |  |  |  |
| 1600 | 1760 |  |  |  |
| 1700 | 1860 |  |  |  |
| 1800 | 1960 |  |  |  |
| 1900 | 2060 |  |  |  |
| 2000 | 2160 |  |  |  |
| 10000 | 10160 |  |  |  |

Table 4: Measured Load Voltage Versus Load Resistance for .

| (Measured) = Volts |
| --- |
| 0 | 160 |  |  |  |
| 100 | 260 |  |  |  |
| 200 | 360 |  |  |  |
| 300 | 460 |  |  |  |
| 400 | 560 |  |  |  |
| 500 | 660 |  |  |  |
| 600 | 760 |  |  |  |
| 700 | 860 |  |  |  |
| 800 | 960 |  |  |  |
| 900 | 1060 |  |  |  |
| 1000 | 1160 |  |  |  |
| 1100 | 1260 |  |  |  |
| 1200 | 1360 |  |  |  |
| 1300 | 1460 |  |  |  |
| 1400 | 1560 |  |  |  |
| 1500 | 1660 |  |  |  |
| 1600 | 1760 |  |  |  |
| 1700 | 1860 |  |  |  |
| 1800 | 1960 |  |  |  |
| 1900 | 2060 |  |  |  |
| 2000 | 2160 |  |  |  |
| 2100 | 2260 |  |  |  |
| 2200 | 2360 |  |  |  |
| 2300 | 2460 |  |  |  |
| 2400 | 2560 |  |  |  |
| 2500 | 2660 |  |  |  |
| 2600 | 2760 |  |  |  |
| 2700 | 2860 |  |  |  |
| 2800 | 2960 |  |  |  |
| 2900 | 3060 |  |  |  |
| 3000 | 3160 |  |  |  |
| 3100 | 3260 |  |  |  |
| 3200 | 3360 |  |  |  |
| 3300 | 3460 |  |  |  |
| 3400 | 3560 |  |  |  |
| 3500 | 3660 |  |  |  |
| 3600 | 3760 |  |  |  |
| 3700 | 3860 |  |  |  |
| 3800 | 3960 |  |  |  |
| 3900 | 4060 |  |  |  |
| 4000 | 4160 |  |  |  |
| 5000 | 5160 |  |  |  |
| 10000 | 10160 |  |  |  |

Table 5: Measured Load Voltage Versus Load Resistance for .

Part C: Photodiode Linearity.

[IMPRTANT: FOR THIS PART OF THE EXPERIMENT, THE LASER MUST BE TURNED ON FOR AT LEAST ONE HOUR BRFORE RELIABLE MEASUREMENTS CAN BE TAKEN].

1- Refer to the experimental setup shown in Figure 7.

2- Set the following circuit setting:  and .

3- Insure that the beam is reasonable centered within the light sensitive area of the diode.

4- Insert a polarizer/analyzer pair between the laser and the photodiode.

5- Set the analyzer angle  to .

6- Rotate the polarizer until  is minimized. The light incident on the photodiode should be too weak to see with the naked eye.

7- Set the analyzer angle  to . Now, the light beam is visible again.

8- Cover the circuit with the carton box in such away that the laser beam passes through the center of the pinhole. [Take care not to disturb the circuit while placing the box].

9- Now, measure  and record its value in table 6.

10- Repeat step 9 for  and all remaining values of  shown in table 6. [Note that the step size for  is not fixed].

11- Remove the circuit and replace it with the optical power meter. Set the meter to dBm and .

12- Adjust the position of the optical power meter for maximum power reading. [DO not remove the cover of the power meter]. There is no need to use the carton box in this case.

13- Set the analyzer angle  to  and record the power meter reading in table 6.

14- Repeat step 13 for  and all the remaining values of  shown in table 6.

15- Calculate the current  and record its value in table 6.

16- Convert the power meter readings to mW and record the values in the fifth column of table 6.

17- Plot  versus the optical power in mW.

18- On the same graph, draw a straight line that best fits the experimental graph, and then comment on the linearity of the photodiode.

19- Why did we use a polarizer/analyzer pair? What function do they serve in this part of the experiment? Why did we use a box cover in the first part? Discuss the experimental results and write some conclusions.

Figure 7: Setup Used for Photodiode Linearity Measurements.

| (Degree) |  | PdBm |  | P (mW) |
| --- | --- | --- | --- | --- |
| 10 |  |  |  |  |
| 20 |  |  |  |  |
| 30 |  |  |  |  |
| 35 |  |  |  |  |
| 40 |  |  |  |  |
| 45 |  |  |  |  |
| 50 |  |  |  |  |
| 55 |  |  |  |  |
| 60 |  |  |  |  |
| 65 |  |  |  |  |
| 70 |  |  |  |  |
| 80 |  |  |  |  |
| 82 |  |  |  |  |
| 84 |  |  |  |  |
| 86 |  |  |  |  |
| 88 |  |  |  |  |
| 90 |  |  |  |  |

Table 6: Experimental Data for Part C.

REPORT REQUIREMENTS:

Part A: The requirements for part A were already stated in the procedure.

Part B:

1- On the same graph, plot  versus  for the cases ,  and .

2- Comment on the maximum voltage obtainable in each case. Compare the results with the theoretical prediction of equation (4).

3- On the same graph, plot  versus  for the cases ,  and .

4- Comment on the behavior of . Is it always constant? If not when is it constant and when is it not constant?

5- On the same graph, plot the photodiode IV characteristics (i.e.  versus , not   versus )  for the cases ,  and .

6- Comment on the resulting graphs. In what way are they different and in what way are they similar?

7- In order to enrich your report, ask yourself the following questions which will help you in developing comments and conclusions appropriate for this part of the experiment.

- Does the photocurrent (increase or decrease) (rapidly or slowly) with the forward bias?

- Do you see approximately straight lines in the IV characteristics?

- Can we model the photodiode as an optical power-dependent current source? If yes, what conditions must be satisfied for the validity of this model?

- Do you expect the IV characteristics of the photodiode to change when the incident optical power changes? Explain.

- Do you expect the IV characteristics of the photodiode to change when the incident light wavelength changes? Explain.

- When , do we obtain a complete or partial IV characteristic? Explain.

PART C: The requirements for part C were already stated in the procedure.

QUESTIONS:

1- Suppose the incident power is extremely low, but measurable, what changes and precautions should be done in part B of this experiment to make it successful?

2- The circuit shown in Figure 4 is much more practical for fiber optics communication than the one shown in Figure 2 (b). Explain why, giving as many reasons as you can think of. [Hint 1: current is harder to deal with than voltage. Hint 2: Different circuits require different voltage levels].

3- When does the circuit used in part B operate in pure photovolatic mode?

4- Using the experimental results, estimate the maximum obtainable voltage  when the photodiode operates in the photovolatic mode.

5- Imagine that you were asked to design a photodetector circuit for use in an optical fiber communication system operating at . Is the photodiode used in this experiment a suitable choice? Justify your answer.

6- In part B of the experiment with , estimate the range of  load resistance  in which we can model the photodiode (approximately) as an ideal current source?

7- In part B of the experiment with  and , use the experimental data to predict the load voltage  when the incident power is reduced by a factor of 2.

8- Answer question (6), assuming the incident power is reduced by a factor of 1000.

9- In part B of the experiment, with  and , the photocurrent  remains almost independent of  in a certain range of . However, with ,  does not remain constant as  changes, even for low values of . Briefly explain why this is so. [Hint: Find out what the bias of the photodiode is when . Is it forward or reverse biased?].

10- Assuming we can model the photodiode as an optical power-dependent current source, develop an appropriate approximate mathematical model for this source [i.e.  , find the approximate value of the constant at ]. Use the experimental data and the known responsivity of the diode for this purpose.

# Text boxes

Incident Light

Incident Light

A

Incident Power

HeNe

Laser

Ammeter

Incident Power

HeNe

Laser

Polarizer

Analyzer

HeNe

Laser